Pumani vs Conventional Bubble CPAP

Summary & observations:

- Pumani & conventional bCPAP devices are different in terms of working principle & construction, a comparative study has been done to identify potential flaws & advantages of the Pumani system.
- During the study it is observed that the pressure stability of the breathing circuit is low & breathing resistance and imposed work of breathing is very high, this can cause serious damage to the patient due to the non-linear characteristics of Flow vs. Pressure(unlike conventional bCPAP).
- It is also claimed that accumulation of CO2 can happen inside the breathing circuit of previous Pumani design.
- In the revised design, bleed valves are introduced to address the CO2 accumulation issue but at the same time it also adds complexity to the circuit & most importantly it adds pressure instability in the breathing circuit.
- With increased flow, the bleed valves exert more pressure to the breathing circuit causing it to work in a non-linear fashion.
- For the revised pumani system, submersion depth could not be used to adjust or control CPAP level without considering fresh gas flow rate in the breathing circuit.