## Sheet1
Breathing circuit pressure validation | [14 empty] | Date: | 26-06-2018
Lattice Innovations
using bubbling humidifier | [14 empty] | New Delhi, India
Results: With surface humidifier | [10 empty] | Results: With bubbling humidifier
10 lpm PEEP | 9 lpm PEEP | 8 lpm PEEP | 7 lpm PEEP | 6 lpm PEEP | 5 lpm PEEP | 4 lpm PEEP | 3 lpm PEEP | 2 lpm PEEP | 1 lpm PEEP | Flow Rate (LPM) | 10 lpm PEEP | 9 lpm PEEP | 8 lpm PEEP | 7 lpm PEEP | 6 lpm PEEP | 5 lpm PEEP | 4 lpm PEEP | 3 lpm PEEP | 2 lpm PEEP | 1 lpm PEEP
9.9 | 9.1 | 7.9 | 7 | 6.1 | 5 | 3.9 | 3 | 2 | 1 | 1 | 10.1 | 9.1 | 8 | 7 | 6.1 | 5 | 3.9 | 3 | 2 | 1
9.9 | 9.1 | 8 | 7 | 6.2 | 5.1 | 4 | 3.1 | 2.1 | 1.1 | 2 | 10.1 | 9.1 | 8 | 7 | 6.2 | 5.1 | 4 | 3.1 | 2.1 | 1
10 | 9.2 | 8 | 7.1 | 6.2 | 5.1 | 4.1 | 3.2 | 2.1 | 1.2 | 3 | 10.1 | 9.1 | 8 | 7.1 | 6.3 | 5.2 | 4.1 | 3.1 | 2.1 | 1.1
10.2 | 9.3 | 8.1 | 7.2 | 6.3 | 5.2 | 4.2 | 3.3 | 2.3 | 1.3 | 4 | 10.2 | 9.2 | 8.1 | 7.1 | 6.4 | 5.3 | 4.1 | 3.2 | 2.2 | 1.2
10.3 | 9.4 | 8.2 | 7.3 | 6.5 | 5.3 | 4.3 | 3.4 | 2.3 | 1.4 | 5 | 10.4 | 9.3 | 8.2 | 7.2 | 6.4 | 5.3 | 4.3 | 3.3 | 2.3 | 1.3
10.3 | 9.5 | 8.4 | 7.4 | 6.6 | 5.5 | 4.4 | 3.5 | 2.4 | 1.5 | 6 | 10.4 | 9.4 | 8.3 | 7.3 | 6.5 | 5.4 | 4.4 | 3.4 | 2.4 | 1.4
10.5 | 9.6 | 8.4 | 7.5 | 6.7 | 5.6 | 4.5 | 3.6 | 2.6 | 1.7 | 7 | 10.6 | 9.5 | 8.4 | 7.4 | 6.6 | 5.6 | 4.5 | 3.5 | 2.6 | 1.5
10.5 | 9.6 | 8.5 | 7.6 | 6.8 | 5.7 | 4.6 | 3.7 | 2.7 | 1.7 | 8 | 10.7 | 9.6 | 8.5 | 7.5 | 6.7 | 5.6 | 4.6 | 3.6 | 2.6 | 1.6
10.6 | 9.7 | 8.6 | 7.7 | 6.9 | 5.8 | 4.8 | 3.8 | 2.8 | 1.8 | 9 | 10.7 | 9.7 | 8.6 | 7.6 | 6.8 | 5.7 | 4.7 | 3.7 | 2.8 | 1.7
10.7 | 9.9 | 8.7 | 7.8 | 7 | 5.9 | 4.9 | 3.9 | 2.9 | 1.9 | 10 | 10.8 | 9.8 | 8.7 | 7.7 | 6.9 | 5.8 | 4.8 | 3.8 | 2.9 | 1.9
Max Deviation: Surface Humidifier | [10 empty] | Max Deviation: Bubbling Humidifier
10 lpm PEEP | 9 lpm PEEP | 8 lpm PEEP | 7 lpm PEEP | 6 lpm PEEP | 5 lpm PEEP | 4 lpm PEEP | 3 lpm PEEP | 2 lpm PEEP | 1 lpm PEEP |  | 10 lpm PEEP | 9 lpm PEEP | 8 lpm PEEP | 7 lpm PEEP | 6 lpm PEEP | 5 lpm PEEP | 4 lpm PEEP | 3 lpm PEEP | 2 lpm PEEP | 1 lpm PEEP
0.8 | 0.8 | 0.8 | 0.8 | 0.9 | 0.9 | 1 | 0.9 | 0.9 | 0.9 |  | 0.7 | 0.7 | 0.7 | 0.7 | 0.8 | 0.8 | 0.9 | 0.8 | 0.9 | 0.9
Difference in breathing circuit air pressure
|  |  | LPM | S.H | B.H
|  |  | 1 | 0.9 | 0.9
|  |  | 2 | 0.9 | 0.9
|  |  | 3 | 0.9 | 0.8
|  |  | 4 | 1 | 0.9
|  |  | 5 | 0.9 | 0.8
|  |  | 6 | 0.9 | 0.8
|  |  | 7 | 0.8 | 0.7
|  |  | 8 | 0.8 | 0.7
|  |  | 9 | 0.8 | 0.7
|  |  | 10 | 0.8 | 0.7